# Testing and test results: Modular tether less baby simulator

Name: 		Hussam Al-Anesi (635155), Stijn Jans (638172), Alif Widianto (642546)

Academy: 	HAN (Hogeschool van Arnhem en Nijmegen)

Education:	Electrical Engineering (IPS), 3rd grade

Location: 	Arnhem

Version:	1.0

Date:		03-06-2022
Pages:		15

# Document history

| Version | Date | Author | Modification |
| --- | --- | --- | --- |
| 1.0 | 03-06-2022 | Hussam Al-Anesi, Stijn Jans, Alif Widianto | All test added |

Table of Contents

Testing and test results: Modular tether less baby simulator	1

Document history	2

Verification of equipment/instrumentation list	4

Testing of Equipment/Instrumentation	6

Test 01.01: BMS Cell balancing	6

Test 01.02: Output voltage of BMS	8

Test 01.03: Simulation of overvoltage condition of BMS	9

Test 01.04: Simulation of a hot and cold condition of BMS	10

Test 01.05: Simulation of undervoltage of BMS	11

Test 02.01: 3.3V and 5V output buck converter	12

Test 03.01: Reading the state of charge using the fuel gauge	13

Testing of the complete system	15

Test 05.01: Complete system	15

# Figure list

Figure 1: BMS cell balancing cell 1	6

Figure 2: BMS cell balancing cell 2	7

Figure 3: BMS cell balancing cell 3	7

Figure 4: Output voltage of BMS	8

Figure 5: Simulation of overvoltage condition of BMS	9

Figure 6: Simulation of a hot and cold condition of BMS	10

Figure 7: Simulation of undervoltage of BMS	11

Figure 8: 3.3V and 5V output buck converter	12

Figure 9: Reading the state of charge using the fuel gauge test setup	13

Figure 10: Reading the state of charge using the fuel gauge software	14

Figure 11: Complete system	15

Figure 12: Complete system without charger	16

Figure 13: Complete system with charger	16

# Verification of equipment/instrumentation list

| Description: |  |  |
| --- | --- | --- |
| Verification of the test setup of the modular tether less baby patient simulator. Verify the installation of all components and equipment. |  |  |
| Acceptance Criteria: |  |  |
| The specifications of the installed equipment/instrumentation should correspond with the tabled data. |  |  |
| Result / Deviations: |  |  |
| Charger was not delivered on time so instead a power supply is used. |  |  |
| Conclusion: | PASS ✓ | FAIL |
| Comments: - |  |  |

| Name / Description | Brand / Draw. No | Specifications | Pass / Fail |
| --- | --- | --- | --- |
| Charger | - | 12V DC adapter 4A | Fail X |
| Batteries | Panasonic NCR18650B ZLF | Technology: Li-ion Voltage: 3.7V Capacity: 3400 mAh | PASS ✓ |
| BMS | Texas Instruments BQ779915EVM-014 | 3-series to 5-series stackable ultra-low power primary protector evaluation module | PASS ✓ |
| Buck converter | Texas Instruments LM2596 | Input: 1.2V – 37V Output: 3.3V, 5V or 12V | PASS ✓ |
| Fuel Gauge | Texas Instruments BQ34Z100EVM | 1s to 16s Impedance Track Fuel Gauge Battery Evaluation Module ¦ Gas Gauge Battery | PASS ✓ |

# Testing of Equipment/Instrumentation

## Test 01.01: BMS Cell balancing

| Step – by – step plan |  |  |
| --- | --- | --- |
| Equipment / Instrumentation reference: |  |  |
| Multimeter |  |  |
| Extra information |  |  |
| Since 3 cells are used, terminals 3, 4 and 5 are shorted. The batteries are connected between terminal bat – and 1, 1 and 2, 2 and 3. |  |  |
| Step 1 |  |  |
| Make sure the batteries are charged through the BMS. |  |  |
| Step 2 |  |  |
| Set the multimeter to voltage measuring range. |  |  |
| Step 3 |  |  |
| Connect the positive terminal of the multimeter to the positive terminal of battery 1. |  |  |
| Step 4 |  |  |
| Connect the negative terminal of the multimeter to the negative terminal of battery 1. |  |  |
| Step 5 |  |  |
| Read the voltage of battery 1 on the display of the multimeter. |  |  |
| Step 6 |  |  |
| Read the voltage of battery 1 on the display of the multimeter. |  |  |
| Step 7 |  |  |
| Repeat step 1 till 6 for battery 2 and 3. |  |  |
| Step 8 |  |  |
| Compare the measured voltages and see if they are approximately equal. |  |  |
| Expected test result: | Measured test result |  |
| Approximately equal voltages | Approximately equal voltages |  |
| Test conclusion | PASS ✓ | FAIL |
| The measured results matched the expected results as can be seen in Figure 1: BMS cell balancing cell 1, Figure 2: BMS cell balancing cell 2, Figure 3: BMS cell balancing cell 3. Figure 1: BMS cell balancing cell 1 Figure 2: BMS cell balancing cell 2 Figure 3: BMS cell balancing cell 3 |  |  |

## Test 01.02: Output voltage of BMS

| Step – by – step plan |  |  |
| --- | --- | --- |
| Equipment / Instrumentation reference: |  |  |
| Power supply, Multimeter |  |  |
| Extra information |  |  |
| For this test use the BQ77915EVM-014 user guide section 2.4 steps 1 till 8: https://www.ti.com/lit/ug/sluubu2b/sluubu2b.pdf?ts=1654166305192&ref_url=https%253A%252F%252Fwww.ti.com%252Ftool%252FBQ77915EVM-014 |  |  |
| Expected test result: | Measured test result |  |
| Output voltage should meet input voltage of about 18V | Output voltage meets the input voltage of about 18V. |  |
| Test conclusion | PASS ✓ | FAIL |
| The measured results matched the expected results as can be seen in Figure 4: Output voltage of BMS Figure 4: Output voltage of BMS |  |  |

## Test 01.03: Simulation of overvoltage condition of BMS

| Step – by – step plan |  |  |
| --- | --- | --- |
| Equipment / Instrumentation reference: |  |  |
| Power supply, Multimeter |  |  |
| Extra information |  |  |
| For this test use the BQ77915EVM-014 user guide section 2.4 steps 9 till 11: https://www.ti.com/lit/ug/sluubu2b/sluubu2b.pdf?ts=1654166305192&ref_url=https%253A%252F%252Fwww.ti.com%252Ftool%252FBQ77915EVM-014 |  |  |
| Expected test result: | Measured test result |  |
| When a supply voltage of 22V is adjusted the voltage over the pack terminals is approximately 600mV below the supply voltage. | The output is approximately 600mV below the supply voltage |  |
| Test conclusion | PASS ✓ | FAIL |
| The measured results matched the expected results as can be seen in Figure 5: Simulation of overvoltage condition of BMS. Figure 5: Simulation of overvoltage condition of BMS |  |  |

## Test 01.04: Simulation of a hot and cold condition of BMS

| Step – by – step plan |  |  |
| --- | --- | --- |
| Equipment / Instrumentation reference: |  |  |
| Power supply, Multimeter |  |  |
| Extra information |  |  |
| For this test use the BQ77915EVM-014 user guide here section 2.4 steps 12 and 13: https://www.ti.com/lit/ug/sluubu2b/sluubu2b.pdf?ts=1654166305192&ref_url=https%253A%252F%252Fwww.ti.com%252Ftool%252FBQ77915EVM-014 |  |  |
| Expected test result: | Measured test result |  |
| When the shunt on pins 1-2 of J4 are removed, the pack voltage drops to approximately 0V. When the shunt on pins 3-4 of J4 are removed, the pack voltage drops to approximately 0V. | The pack voltage dropped during both measurements to approximately 0V. |  |
| Test conclusion | PASS ✓ | FAIL |
| The measured results matched the expected results as can be seen in Figure 6: Simulation of a hot and cold condition of BMS. Figure 6: Simulation of a hot and cold condition of BMS |  |  |

## Test 01.05: Simulation of undervoltage of BMS

| Step – by – step plan |  |  |
| --- | --- | --- |
| Equipment / Instrumentation reference: |  |  |
| Power supply, Multimeter |  |  |
| Extra information |  |  |
| For this test use the BQ77915EVM-014 user guide section 2.4 steps 14 till 16: https://www.ti.com/lit/ug/sluubu2b/sluubu2b.pdf?ts=1654166305192&ref_url=https%253A%252F%252Fwww.ti.com%252Ftool%252FBQ77915EVM-014 |  |  |
| Expected test result: | Measured test result |  |
| When a supply voltage of 12V is adjusted the voltage over the pack terminals drops to approximately 0V. | The voltage over the pack terminals is approximately 0V. |  |
| Test conclusion | PASS ✓ | FAIL |
| The measured results matched the expected results as can be seen in Figure 7: Simulation of undervoltage of BMS. Figure 7: Simulation of undervoltage of BMS |  |  |

## Test 02.01: 3.3V and 5V output buck converter

| Step – by – step plan |  |  |
| --- | --- | --- |
| Equipment / Instrumentation reference: |  |  |
| Power supply, Multimeter |  |  |
| Extra information |  |  |
| The potentiometer of the buck converters are adjusted to achieve 3.3V and 5V output. |  |  |
| Step 1 |  |  |
| Connect the positive terminal of the power supply to the positive terminal of the buck converter. |  |  |
| Step 2 |  |  |
| Connect the negative terminal of the power supply to the negative terminal of the buck converter. |  |  |
| Step 3 |  |  |
| Connect the positive terminal of the multimeter to the positive terminal of the buck converter. |  |  |
| Step 4 |  |  |
| Connect the negative terminal of the multimeter to the negative terminal of the buck converter. |  |  |
| Step 5 |  |  |
| Set the multimeter to voltage measuring range. |  |  |
| Step 6 |  |  |
| Set the power supply to 12V and turn it on. |  |  |
| Expected test result: | Measured test result |  |
| Measurement of about 3.3V and 5 V at the multimeter. | Output voltage measurement meets the voltages 3.3V and 5V. |  |
| Test conclusion | PASS ✓ | FAIL |
| The measured results matched the expected results as can be seen in Figure 8: 3.3V and 5V output buck converter Figure 8: 3.3V and 5V output buck converter |  |  |

## Test 03.01: Reading the state of charge using the fuel gauge

| Step – by – step plan |  |  |
| --- | --- | --- |
| Equipment / Instrumentation reference: |  |  |
| BQ34Z100 Fuel Gauge Evaluation Board from Texas Instruments (fuel gauge EVM) EV2400 Analog Evaluation Model from Texas Instruments (Fuel gauge adapter) Bq Battery Management Software from Texas Instruments |  |  |
| Extra information |  |  |
| The steps on conducting the test and experimentation can be followed from these datasheets: The guide for connection between laptop with the analog evaluation model can be found here (section 3, Page 2) Quickstart Guide for the bq34z100-G1 (ti.com) The steps to configure the fuel gauge EVM can be found here ( Section 2.3, Page 5) bq34z100EVM Wide Range Impedance Track Enabled Battery Fuel Gauge (Rev. B) (ti.com) The bq Battery Management Software can be found here BQSTUDIO Application software & framework ¦ TI.com |  |  |
| Expected test result: | Measured test result |  |
| State of charge of the battery should be read by the software while the battery is connected to the fuel gauge. | The system can properly read the state of charge, state of health, temperature, etc…, of the battery with a single cell. |  |
| Test conclusion | PASS ✓ | FAIL |
| The state of health value is read as expected when one cell is connected to the BMS; however, with multi-cell batteries, some calibration is required on the software side. The measured results matched the expected results as can be seen in Figure 9: Reading the state of charge using the fuel gauge test setup, Figure 10: Reading the state of charge using the fuel gauge software. Figure 9: Reading the state of charge using the fuel gauge test setup Figure 10: Reading the state of charge using the fuel gauge software |  |  |

# Testing of the complete system

## Test 05.01: Complete system

| Step – by – step plan |  |  |
| --- | --- | --- |
| Equipment / Instrumentation reference: |  |  |
| Power supply, Multimeter, 2x 100 Ohm load |  |  |
| Extra information |  |  |
| Connect the system as shown in Figure 11: Complete system Figure 11: Complete system |  |  |
| Step 1 |  |  |
| Measure the voltage over the 3.3V load and 5V load without charger connected. |  |  |
| Step 2 |  |  |
| Measure the voltage over the 3.3V load and 5V load with charger connected. |  |  |
| Step 3 |  |  |
| Read the state of charge using the fuel Gauge. |  |  |
| Expected test result: | Measured test result |  |
| Output of the 3.3V buck converter without charger connected is approximately 3.3V | Measured value is 3.33V |  |
| Output of the 5V buck converter without charger connected is approximately 5V. | Measured value is 5.009V |  |
| Output of the 3.3V buck converter with charger connected is approximately 3.3V | Measured value is 3.33V. |  |
| Output of the 5V buck converter with charger connected is approximately 5V. | Measured value is 5.009V. |  |
| The state of charge should be measured by the fuel gauge reading one cell. | The software couldn’t communicate to the fuel gauge. |  |
| Test conclusion | PASS | FAIL X |
| The measured results matched the expected results as can be seen in Figure 12: Complete system without charger, Figure 13: Complete system with charger. Figure 12: Complete system without charger Figure 13: Complete system with charger |  |  |